TEST REPORT: 7191064517-CHM13-02D-TSL

| Date: | 12 JUL 2013 | Tel: +65 68851312 Fax: +65 67784301 |
| --- | --- | --- |
| Client’s Ref: 221402143Rev1 | Email: zhou.xiao@tuv-sud-psb.sg |

Note: This report is issued subject to the Testing and Certification Regulations of the TÜV SÜD Group and the General Terms and Conditions of Business of TÜV SÜD PSB Pte Ltd. In addition, this report is governed by the terms set out within this report.

SUBJECT

Overall Migration Test for “Nitrile Blue Powder Free Finger Textured (NBPFFT) ENLW35” Sample

SAMPLE SUBMISSION DATE

26 Jun 2013

DESCRIPTION OF SAMPLE

One packet of glove sample labelled as follows was received. The test was confirmed to be analysed on 01 Jul 2013.

| Product Name | Quantity |
| --- | --- |
| Nitrile Blue Powder Free Finger Textured (NBPFFT) ENLW35 | 25pcs |

DATE OF ANALYSIS

03 Jul 2013 –  12 Jul 2013

| Laboratory: | Phone : +65-6885 1333 | Regional Head Office: |
| --- | --- | --- |
| TÜV SÜD PSB Pte. Ltd. | Fax : +65-6776 8670 | TÜV SÜD Asia Pacific Pte. Ltd. |
| No.1 Science Park Drive | E-mail: testing@tuv-sud-psb.sg | 3 Science Park Drive, #04-01/05 |
| Singapore 118221 | www.tuv-sud-psb.sg | The Franklin, Singapore 118223 |
|  | Co. Reg : 199002667R | Page 1 of 3 |

TEST REPORT: 7191064517-CHM13-02D-TSL

12 JUL 2013

METHOD OF TEST

The sample was analysed for the following tests according to Commission Regulation (EU) No 10/2011:

- Preparation of Test Specimen

Only the exterior of the glove sample was performed for the test.

- Overall Migration Content with Aqueous Food Simulant (50% Ethanol)

According to BS EN 1186-9:2002 – Test Methods for overall migration into aqueous food simulants

by article filling.

RESULTS

Table 1 : Overall Migration Content with Food Simulant for the “Nitrile Blue Powder Free Finger Textured (NBPFFT) ENLW35” Sample

|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Commission |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Testing |  |  | Surface |  |  | Volume of |  |  | Overall |  |  | Regulation (EU) |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | No 10/2011 |  |  |
|  | Type of Simulant |  |  |  |  |  | Area |  |  | Extractant |  |  | Migration |  |  |  |  |  |
|  |  |  |  | Condition |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Requirement for |  |  |
|  |  |  |  |  |  |  | 2 |  |  |  |  |  | 2 |  |  |  |  |  |
|  |  |  |  |  |  |  | (dm ) |  |  | (ml) |  |  | (mg/dm ) |  |  | Overall Migration |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Content (mg/dm2) |  |  |
|  | 1. 50% Ethanol |  | 40 °C, 2 hours | 5.01 |  | 250 |  | 2.5 |  | <10 |  |  |

Based on the above result, the “Nitrile Blue Powder Free Finger Textured (NBPFFT) ENLW35” sample meets the overall migration requirement under Commission Regulation (EU) No 10/2011 – “Plastic materials and articles shall not transfer their constituents to foodstuffs in quantities exceeding 10 milligrams of total constituents released per dm2 of food contact surface (mg/dm2) (overall migration limit)”.

MS TAN SER LING	Dr XIAO ZHOU

TECHNICAL EXECUTIVE	PRODUCT MANAGER

MICROCONTAMINATION DIAGNOSIS

CHEMICAL & MATERIALS

Page 2 of 3

TEST REPORT: 7191064517-CHM13-02D-TSL

12 JUL 2013

Please note that this Report is issued under the following terms :

- This report applies to the sample of the specific product/equipment given at the time of its testing/calibration. The results are not used to indicate or imply that they are applicable to other similar items. In addition, such results must not be used to indicate or imply that TÜV SÜD PSB approves, recommends or endorses the manufacturer, supplier or user of such product/equipment, or that TÜV SÜD PSB in any way “guarantees” the later performance of the product/equipment. Unless otherwise stated in this report, no tests were conducted to determine long term effects of using the specific product/equipment.

- The sample/s mentioned in this report is/are submitted/supplied/manufactured by the Client. TÜV SÜD PSB therefore assumes no responsibility for the accuracy of information on the brand name, model number, origin of manufacture, consignment or any information supplied.

- Nothing in this report shall be interpreted to mean that TÜV SÜD PSB has verified or ascertained any endorsement or marks from any other testing authority or bodies that may be found on that sample.

- This report shall not be reproduced wholly or in parts and no reference shall be made by the Client to TÜV SÜD PSB or to the report or results furnished by TÜV SÜD PSB in any advertisements or sales promotion.

- Unless otherwise stated, the tests were carried out in TÜV SÜD PSB Pte Ltd, No.1 Science Park Drive Singapore 118221.

July 2011

Page 3 of 3